|  | Certificate Number: UMECC xxx /xxxx ( xx pages) Performance Test Certificate for Patient Monitor شهاده اختبار اداء وظيفى لجهاز شاشة مراقبة المريض |  |
| --- | --- | --- |

بيانات المعمل واجهزة القياس الامامى المستخدمة

| {TestLab} |  |  | معمل الاختبار | بيانات المعمل |
| --- | --- | --- | --- | --- |
| {EngineerName} | القائم بالأختبار | {TestType} | نوع الاختبار | بيانات الاختبار |
| {TestDeviceMfr} | ماركة الجهاز | {TestDeviceManufacturer} | الجهاز المستخدم في الاختبار | بيانات جهاز الأختبار |
| {TestDeviceModel} | موديل الجهاز |  |  |  |
| {TestDeviceSerialNo} | رقم مسلسل |  |  |  |
| {VisitDate} | تاريخ انتهاء معايرة الجهاز | {OrderDate} | تاريخ معايرة الجهاز |  |

بيانات الجهاز الطبى المختبر

| بيانات مكان الجهاز الطبى المختبر | اسم المستشفى المختبر بها الجهاز الطبى | {HospitalName} |  |  |
| --- | --- | --- | --- | --- |
| بيانات الجهاز الطبى المختبر | اسم الجهاز | Patient Monitor | ماركة | {Manufacturer} |
|  |  | شاشة مراقبة مريض | موديل | {Model} |
|  |  |  | رقم مسلسل | {SerialNo} |

نتيجة الاختبار

| {Final_Qualitative} |  | Qualitative Test | {Final} |  | نتيجة الاختبار |
| --- | --- | --- | --- | --- | --- |
| {Final_Quantitative} |  | Quantitative Test |  |  |  |
| {VisitDate} |  |  |  |  | تاريخ الاختبار |
| رئيس مركز الاستشارات الهندسية وعميد كلية الهندسة |  | المدير الفنى لوحدة يوميك |  |  | القائم بالمعايرة |  |

Qualitative Test

| Visual Inspection |  |  |  |
| --- | --- | --- | --- |
| Item | Status | Item | Status |
| Chassis/Housing | {Cha} | Controls /Switches | {Con} |
| Mount | {Mou} | Battery/charger | {Bat} |
| Casters/Brakes | {Cas} | Indicator/Displays | {Ind} |
| AC plug | {AC} | Labeling | {Lab} |
| Line Cord | {Lin} | Alarms | {Ala} |
| Screen | {Scr} | Module Housing | {Hou} |
| SPO2 cable | {SPO} | Mounting/Trolley | {Tro} |
| Representation capability of ECG |  |  |  |
| Atrial Fibrillation | {Atr} | Premature ventricle contraction | {Prem} |
| Ventricle Fibrillation | {Ven} | Paroxysmal Atrial Tachycardia (PAT) | {Paro} |
| Atrial Flutter | {Atr} | Polymorphic Ventricular Tachycardia (PVT) | {Poly} |
| Representation of Standard signals (Triangle, Square, Sinusoid) | {Rep} | Represent ECG waveforms with different Amplitudes 0.5, 1,1.5,2,2.5,3,3.5 | {ECG} |

Quantitative Test

| Heart Rate Measurement Based on NORMAL ECG RHYTHM |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Setting value (BPM*) | Average Measured value* | Error** | Accepted Range ± 5% ± 1*** | Status | Uncertainty (Type A) |  |
| {Heart_set40} | {Heart_Ave40} | {Heart_Error40} | {Heart_Acc40} | {Heart_Sta40} | {Heart_unc40} |  |
| {Heart_set60} | {Heart_Ave60} | {Heart_Error60} | {Heart_Acc60} | {Heart_Sta60} | {Heart_unc60} |  |
| {Heart_set80} | {Heart_Ave80} | {Heart_Error80} | {Heart_Acc80} | {Heart_Sta80} | {Heart_unc80} |  |
| {Heart_set100} | {Heart_Ave100} | {Heart_Error100} | {Heart_Acc100} | {Heart_Sta100} | {Heart_unc100} |  |
| {Heart_set150} | {Heart_Ave150} | {Heart_Error150} | {Heart_Acc150} | {Heart_Sta150} | {Heart_unc150} |  |
| {Heart_set200} | {Heart_Ave200} | {Heart_Error200} | {Heart_Acc200} | {Heart_Sta200} | {Heart_unc200} |  |
| *Average measured value calculated for at least three readings ** Error is the difference between setting value and Average measured value. *** The limitation is calculated be ± 5 % nominal heart rate accuracy−/+ 1 BPM accuracy of the used patient monitor simulator. |  |  |  |  |  |  |
| رئيس مركز الاستشارات الهندسية وعميد كلية الهندسة |  | المدير الفنى لوحدة يوميك |  | القائم بالمعايرة |  |
| ا د مصطفى محمود |  | د ايناس اسماعيل محمد |  | م /-------- |  |

(UMECC xx/xxxx: Page2/xx)

| Respiration Measurement Based on ECG Cable |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Setting value (BPM*) | Average Measured value** | Error*** | Accepted Range ± 1%**** | Status | Uncertainty (Type A) |
| {Resp_Set5} | {Resp_Avg5} | {Resp_Err5} | {Resp_Acc5} | {Resp_Sta5} | {Resp_Unc5} |
| {Resp_Set10} | {Resp_Avg10} | {Resp_Err10} | {Resp_Acc10} | {Resp_Sta10} | {Resp_Unc10} |
| {Resp_Set15} | {Resp_Avg15} | {Resp_Err15} | {Resp_Acc15} | {Resp_Sta15} | {Resp_Unc15} |
| {Resp_Set30} | {Resp_Avg30} | {Resp_Err30} | {Resp_Acc30} | {Resp_Sta30} | {Resp_Unc30} |
| *Breaths Per Minute **Average measured value calculated for at least three readings *** Error is the difference between setting value and Average measured value. **** The limitation is calculated be ± 5% nominal RR level -+4% accuracy of the used patient monitor simulator. |  |  |  |  |  |

| NIBP Measurement |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Setting value (mmHg) | Average Measured value* |  | Error** | Accepted Range*** ± 10 mmHg-0.5% |  | Status |  | Uncertainty (Type A) |  |
| {Set_Sys60} | {NIBP_S_Avg60} |  | {NIBP_S_Err60} | {NIBP_S_Acc60} |  | {NIBP_S_Sta60} |  | {NIBP_S_Unc60} |  |
| {Set_Dia30} | {NIBP_D_Avg30} |  | {NIBP_D_Err30} | {NIBP_D_Acc30} |  | {NIBP_D_Sta30} |  | {NIBP_D_Unc30} |  |
| {Set_Sys80} | {NIBP_S_Avg80} |  | {NIBP_S_Err80} | {NIBP_S_Acc80} |  | {NIBP_S_Sta80} |  | {NIBP_S_Unc80} |  |
| {Set_Dia40} | {NIBP_D_Avg40} |  | {NIBP_D_Err40} | {NIBP_D_Acc40} |  | {NIBP_D_Sta40} |  | {NIBP_D_Unc40} |  |
| {Set_Sys100} | {NIBP_S_Avg100} |  | {NIBP_S_Err100} | {NIBP_S_Acc100} |  | {NIBP_S_Sta100} |  | {NIBP_S_Unc100} |  |
| {Set_Dia60} | {NIBP_D_Avg60} |  | {NIBP_D_Err60} | {NIBP_D_Acc60} |  | {NIBP_D_Sta60} |  | {NIBP_D_Unc60} |  |
| {Set_Sys120} | {NIBP_S_Avg120} |  | {NIBP_S_Err120} | {NIBP_S_Acc120} |  | {NIBP_S_Sta120} |  | {NIBP_S_Unc120} |  |
| {Set_Dia80} | {NIBP_D_Avg80} |  | {NIBP_D_Err80} | {NIBP_D_Acc80} |  | {NIBP_D_Sta80} |  | {NIBP_D_Unc80} |  |
| {Set_Sys180} | {NIBP_S_Avg180} |  | {NIBP_S_Err180} | {NIBP_S_Acc180} |  | {NIBP_S_Sta180} |  | {NIBP_S_Unc180} |  |
| {Set_Dia140} | {NIBP_D_Avg140} |  | {NIBP_D_Err140} | {NIBP_D_Acc140} |  | {NIBP_D_Sta140} |  | {NIBP_D_Unc140} |  |
| *Average measured value calculated for at least three readings ** Error is the difference between setting value and Average measured value. *** The limitation is calculated be ±10 nominal NIBP accuracy ±10 mmhg−0.5% accuracy of the used patient monitor simulator. |  |  |  |  |  |  |  |  |  |
| رئيس مركز الاستشارات الهندسية وعميد كلية الهندسة |  |  | المدير الفنى لوحدة يوميك |  | القائم بالمعايرة |  |
| ا د مصطفى محمود |  |  | د ايناس اسماعيل محمد |  | م/ --------- |  |

(UMECC xx/xxxx: Page3/xx)

| SPO2 Measurement |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Setting value (%) | Average Measured value* | Error** | Accepted Range ±2%*** | Status | Uncertainty (Type A) |
| {SPO2_Set75} | {SPO2_Avg75} | {SPO2_Err75} | {SPO2_Acc75} | {SPO2_Sta75} | {SPO2_Unc75} |
| {SPO2_Set85} | {SPO2_Avg85} | {SPO2_Err85} | {SPO2_Acc85} | {SPO2_Sta85} | {SPO2_Unc85} |
| {SPO2_Set90} | {SPO2_Avg90} | {SPO2_Err90} | {SPO2_Acc90} | {SPO2_Sta90} | {SPO2_Unc90} |
| {SPO2_Set94} | {SPO2_Avg94} | {SPO2_Err94} | {SPO2_Acc94} | {SPO2_Sta94} | {SPO2_Unc94} |
| {SPO2_Set98} | {SPO2_Avg98} | {SPO2_Err98} | {SPO2_Acc98} | {SPO2_Sta98} | {SPO2_Unc98} |
| *Average measured value calculated for at least three readings ** Error is the difference between setting value and Average measured value. *** The limitation is calculated be ±3% nominal SPO2% level ±1% accuracy of the used Spo2 Simulator. |  |  |  |  |  |

| Temperature Measurement (MODULE 1) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Setting value (οC) | Average Measured value* | Error** | Accepted Range ± 0.2%*** | Status | Uncertainty (Type A) |
| {Tem_Set33} | {Tem_Avg33} | {Tem_Err33} | {Tem_Acc33} | {Tem_Sta33} | {Tem_Unc33} |
| {Tem_Set37} | {Tem_Avg37} | {Tem_Err37} | {Tem_Acc37} | {Tem_Sta37} | {Tem_Unc37} |
| {Tem_Set41} | {Tem_Avg41} | {Tem_Err41} | {Tem_Acc41} | {Tem_Sta41} | {Tem_Unc41} |
| *Average measured value calculated for at least three readings ** Error is the difference between setting value and Average measured value. *** The limitation is calculated be ± 0.3 οC nominal Temperature level -+0.1% accuracy of the used patient monitor simulator. |  |  |  |  |  |

| Temperature Measurement (MODULE 2) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Setting value (οC) |  | Average Measured value* | Error** |  | Accepted Range ± 0.2%*** | Status |  | Uncertainty (Type A) |
| {Tem2_Set33} |  | {Tem2_Avg33} | {Tem2_Err33} |  | {Tem2_Acc33} | {Tem2_Sta33} |  | {Tem2_Unc33} |
| {Tem2_Set37} |  | {Tem2_Avg37} | {Tem2_Err37} |  | {Tem2_Acc37} | {Tem2_Sta37} |  | {Tem2_Unc37} |
| {Tem2_Set41} |  | {Tem2_Avg41} | {Tem2_Err41} |  | {Tem2_Acc41} | {Tem2_Sta41} |  | {Tem2_Unc41} |
| *Average measured value calculated for at least three readings ** Error is the difference between setting value and Average measured value. *** The limitation is calculated be ± 0.3 οC nominal Temperature level -+0.1% accuracy of the used patient monitor simulator. |  |  |  |  |  |  |  |  |
| رئيس مركز الاستشارات الهندسية وعميد كلية الهندسة |  |  | المدير الفنى لوحدة يوميك |  |  | القائم بالمعايرة |  |
| ا د مصطفى محمود |  |  | د ايناس اسماعيل محمد |  |  |  |  |

(UMECC xx/xxxx: Page4/xx)