|  | Certificate Number: UMECC 384/2026 (3 pages) Performance Test Certificate for SYRINGE PUMP شهاده اختبار اداء وظيفى لجهاز سرنجة محاليل |  |
| --- | --- | --- |

بيانات المعمل واجهزة القياس الامامى المستخدمة

| {TestLab} |  |  | معمل الاختبار | بيانات المعمل |
| --- | --- | --- | --- | --- |
| {EngineerName} | القائم بالأختبار | {TestType} | نوع الاختبار | بيانات الاختبار |
| {TestDeviceMfr} | ماركة الجهاز | {TestDeviceManufacturer} | الجهاز المستخدم في الاختبار | بيانات جهاز الأختبار |
| {TestDeviceModel} | موديل الجهاز |  |  |  |
| {TestDeviceSerialNo} | رقم مسلسل |  |  |  |
| {VisitDate} | تاريخ انتهاء معايرة الجهاز | {OrderDate} | تاريخ معايرة الجهاز |  |

بيانات الجهاز الطبى المختبر

| بيانات مكان الجهاز الطبى المختبر | اسم المستشفى المختبر بها الجهاز الطبى | {HospitalName} |  |  |
| --- | --- | --- | --- | --- |
| بيانات الجهاز الطبى المختبر | اسم الجهاز | SYRINGE PUMP | ماركة | {Manufacturer} |
|  |  | سرنجة محاليل | موديل | {Model} |
|  |  |  | رقم مسلسل | {SerialNo} |

نتيجة الاختبار

| {Final_Qualitative} | Qualitative Test | {Final} | نتيجة الاختبار |
| --- | --- | --- | --- |
| {Final_Quantitative} | Quantitative Test |  |  |
| {VisitDate} |  |  | تاريخ الاختبار |

Qualitative Test

| Visual Inspection |  |  |  |
| --- | --- | --- | --- |
| Item | Status | Item | Status |
| Chassis/Housing | {Cha} | Controls /Switches | {Con} |
| Mount | {Mou} | Door/Misloaded Infusion Set | {Doo} |
| Casters/Brakes | {cas} | Battery/charger | {Bat} |
| AC plug | {AC} | Indicator/Displays | {Ind} |
| Line Cord | {Lin} | Labeling | {Lab} |
| Cables | {Cab} | Air-in-Line | {Air} |
| Screen | {Scr} | Empty Container | {Emp} |
| Flow-Stop Mechanism(s) | {Flo} | Infusion Complete | {Inf} |

Quantitative Test

| - Flow rate Measurement |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Setting value (mL\hr*) | Average Measured value** | Error*** | Accepted Range ± 9%**** | Status | Uncertainty (Type A) |
| {flow_set_10} | {flow_avg_10} | {flow_err_10} | {flow_acc_10} | {flow_sta_10} | {flow_unc_10} |
| {flow_set_15} | {flow_avg_15} | {flow_err_15} | {flow_acc_15} | {flow_sta_15} | {flow_unc_15} |
| {flow_set_20} | {flow_avg_20} | {flow_err_20} | {flow_acc_20} | {flow_sta_20} | {flow_unc_20} |
| *mL\hr =milliliter per hour **Average measured value calculated for at least three readings. *** Error is the difference between setting value and Average measured value. ****The accepted range is calculated to be ± 10 % nominal accepted range ± 1 % the accuracy of the used syringe analyzer based on its manufacture' manual. |  |  |  |  |  |

| رئيس مركز الاستشارات الهندسية وعميد كلية الهندسة | المدير الفنى لوحدة يوميك | القائم بالمعايرة |
| --- | --- | --- |
| ا د مصطفى أحمد الراوي | د ايناس اسماعيل محمد | م حسام ياسر |

(UMECC 384/2026: Page2/3)

| - Occlusion pressure measurement |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Peak value (mmHg) |  | Average Measured value* |  | Accepted Range <723.8** | Status | Uncertainty (Type A) |
|  |  | {occ_avg_peak} |  |  | {occ_sta_peak} | {occ_unc_peak} |
| Time to Alarm (sec) |  | Average Measured value* |  | Accepted Range <12 sec*** | Status | Uncertainty (Type A) |
|  |  | {occ_avg_time} |  |  | {occ_sta_peak} | {occ_unc_peak} |
| *Average measured value calculated for at least three readings. ** The accepted range is calculated to be < 15 psi nominal accepted range ± 1 psi the accuracy of the used syringe analyzer based on its manufacture' manual 723.8 mmHg. *** The accepted range is calculated to be < 12 sec nominal accepted range. |  |  |  |  |  |  |

| رئيس مركز الاستشارات الهندسية وعميد كلية الهندسة | المدير الفنى لوحدة يوميك | القائم بالمعايرة |
| --- | --- | --- |
| ا د مصطفى أحمد الراوي | د ايناس اسماعيل محمد | م حسام ياسر |

(UMECC 384/2026: Page3/3)