|  | Certificate Number: UMECC 148/2025 (3 pages) Performance Test Certificate for Infusion Pump شهاده اختبار اداء وظيفى لجهاز مضخة محاليل |  |
| --- | --- | --- |

بيانات المعمل واجهزة القياس الامامى المستخدمة

| معمل معايرة الأجهزة الطبية التابع لوحدة معايرة واستشارات الأجهزة الطبية (يوميك) التابعة لمركز الاستشارات الهندسية بكلية الهندسة جامعة المنيا |  |  | معمل الاختبار | بيانات المعمل |
| --- | --- | --- | --- | --- |
| د ايناس إسماعيل محمد | القائم بالأختبار | اختبار اداء وظيفى | نوع الاختبار | بيانات الاختبار |
|  |  | لجهاز مضخة محاليل |  |  |
| RIGEL | ماركة الجهاز | جهاز اختبار جهاز مضخة المحاليل | الجهاز المستخدم في الاختبار | بيانات جهاز الأختبار |
| MULTI-FLO | موديل الجهاز |  |  |  |
| 06K-0290 | رقم مسلسل |  |  |  |
| 28/5/2026 | تاريخ انتهاء معايرة الجهاز | 28/5/2025 | تاريخ معايرة الجهاز |  |

بيانات الجهاز الطبى المختبر

| بيانات مكان الجهاز الطبى المختبر | اسم المستشفى المختبر بها الجهاز الطبى | مستشفى الشيخ زايد ال نهيان |  |  |
| --- | --- | --- | --- | --- |
| بيانات الجهاز الطبى المختبر | اسم الجهاز | INFUSION PUMP | ماركة | Sk |
|  |  | مضخة محاليل | موديل | Sk-600i |
|  |  |  | رقم مسلسل | SK81102309 |

نتيجة الاختبار

| PASS | Qualitative Test | Fail | نتيجة الاختبار |
| --- | --- | --- | --- |
| Fail | Quantitative Test |  |  |
| 16-7-2025 |  |  | تاريخ الاختبار |

| رئيس مركز الاستشارات الهندسية وعميد كلية الهندسة | المدير الفنى لوحدة يوميك | القائم بالمعايرة |
| --- | --- | --- |
| ا د مصطفى محمود | د ايناس اسماعيل محمد | د ايناس اسماعيل محمد |

(UMECC 148/2025: Page1/3)

Qualitative Test

| Visual Inspection |  |  |  |
| --- | --- | --- | --- |
| Item | Status | Item | Status |
| Chassis/Housing | Pass | Controls /Switches | Pass |
| Mount | Pass | Door/Misloaded Infusion Set | Pass |
| Casters/Brakes | Pass | Battery/charger | Pass |
| AC plug | Pass | Indicator/Displays | Pass |
| Line Cord | Pass | Labeling | Pass |
| Cables | Pass | Air-in-Line | Pass |
| Screen | Pass | Empty Container | Pass |
| Flow-Stop Mechanism(s) | Pass | Infusion Complete | Pass |

Quantitative Test

| - Flow rate Measurement |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Setting value (mL\hr*) | Average Measured value** | Error*** | Accepted Range ± 9%**** | Status | Uncertainty (Type A) |
| 30 | 33.60 | -3.60 | (27.3 -32.7) | FAIL | 0.85 |
| 60 | 61.16 | -1.16 | (54.6 -65.4) | PASS | 0.69 |
| 100 | 93.52 | 6.48 | (91 -109) | PASS | 2.78 |
| 240 | 210.61 | 29.39 | (218.4 - 261.6) | FAIL | 4.09 |
| 300 | 252.00 | 48.00 | (273 - 327) | FAIL | 0.71 |
| 600 | 463.73 | 136.27 | (546 - 654) | FAIL | 2.03 |
| *mL\hr =milliliter per hour **Average measured value calculated for at least three readings. *** Error is the difference between setting value and Average measured value. ****The accepted range is calculated to be ± 10 % nominal accepted range ± 1 % the accuracy of the used syringe analyzer based on its manufacture' manual. |  |  |  |  |  |

| رئيس مركز الاستشارات الهندسية وعميد كلية الهندسة | المدير الفنى لوحدة يوميك | القائم بالمعايرة |
| --- | --- | --- |
| ا د مصطفى محمود | د ايناس اسماعيل محمد | د ايناس اسماعيل محمد |

(UMECC 148/2025: Page2/3)

| - Occlusion pressure measurement |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Peak value (mmHg) | Average Measured value* |  | Accepted Range <723.8** | Status | Uncertainty (Type A) |
|  | 276.33 |  |  | Pass | 44.95 |
| Time to Alarm (sec) | Average Measured value* |  | Accepted Range <=12 sec*** | Status | Uncertainty (Type A) |
|  | 9 |  |  | PASS | 0.71- |
| *Average measured value calculated for at least three readings. ** The accepted range is calculated to be < 15 psi nominal accepted range ± 1 psi the accuracy of the used syringe analyzer based on its manufacture' manual 723.8 mmHg. *** The accepted range is calculated to be < 12 sec nominal accepted range. |  |  |  |  |  |

| رئيس مركز الاستشارات الهندسية وعميد كلية الهندسة | المدير الفنى لوحدة يوميك | القائم بالمعايرة |
| --- | --- | --- |
| ا د مصطفى محمود | د ايناس اسماعيل محمد | د ايناس اسماعيل محمد |

(UMECC 148/2025: Page3/3)